Colour probe for the visually handicapped persons

Fumiya Furuno

A battery-powered pocketable colour probe is now under developmentwhich

discriminates the colour of an object and using a voice output tells the

persons its colour. Thirteen kinds of colours can be discriminated so far.

These are white, grey, black, red, pink, brown, orange, yellow, green,

blue, light-blue, dark-blue, and purple. This probe is also given a

function of a light probe. Its size, at present, is about 13.8 x q,q x 2.g

cm ' and it weighs about 150 g. It will serve the visually handicapped

persons in their daily life, vocational life and in education. The

developement is now proceeding towards higher accuracy and increasing the

number of colours discriminated.

1. Introduction

The visually handicapped persons can acquire much information of an object

near them by touching it mainly by their fingers and hands. For example,

they can recognize its shape, its material, its temperature, its weight,

its size and roughness of its surface. They, however, can never recognize

the colour of an object at all.We are now developing a small apparatus for

discriminating colour of an object and telling the persons the name of the

colour. The last year,a prototype of the colour probe was presented at the

conference(I) and on the journa1(2).In this workshop an improved colour

probe is presented which has been miniaturized and to which new functions

such as light sensing function have been added.The result of test and

evaluation and that of hearing from the disabled persons who trially used

this probe are also presented.

2. Principle to recognize colours

The colour of an object is determined by both the spectral distribution of

the light thrown onto the object and the spectral reflectance of its

surface. A small lamp placed in the pointed head of this probe radiates the

light on an object and the spectral distribution of the direct light from

the lamp and that of the reflected light from the object are measured. In

order to measure the distribution and the reflectance stated above, six

light sensors are used in this probe.The colour can be calculated out by

putting the output from the six sensors into some formulae experimentally

derived.

3. Principle of the light probe

One of the six sensors in this probe which has the same spectral luminus

efficiency as the normal human eye is used to detect the intensity of the

visual light. The output of this sensor is converted to a frequency of

sound proportional to the intensity of the input light.

4. System configuration

The lamp controlled by the microcomputer illuminates an object and the

direct light from the lamp is received by the three pieces of light sensors

and the reflected light from the object is received by other three light

sensors. The output of these six sensors are amplified,multiplexed,and

converted to digital values and introduced to the microcomputer, which

calculates X, Y and Z called tristimulus values and then x, y called

chromaticity coordinates. AII colours are specified by the above x, y and Y

values. The microcomputer decides the colour from these x,y and Y values

and then selects the word in the ROM corresponding to the decided

colour.The selected word is then picked up and converted to analog

value,amplified and spoken out from the speaker or the earphone. On the use

as the light probe,the lamp is left off and only one sensor's output is

amplified, converted to digital value and introduced to the microcomputer.

The frequency of sound proportional to the intensity of the input light is

produced by the microcomputer and is converted to analog value and then put

out from the speaker or the earphone.

5. Structure and operation

This probe is so designed that it can be carried in a pocket or a handbag

and that it can be operated by one hand.Its size is 13.8x4.4x2.8 cm3 and it

weighs abot 150 g. Fig.l shows the dimensions and the structure. An 8 bit

microcomputer plays an indispensable role to have made this probe small

sized.

When a person wants to know the colour of some object, all he or she has to

do is to put on the power switch,to throw the mode switch to the colour

mode, to put the tip of the probe on the object and to push the push

switch. Then, the voice representing the colour can beheard from the

speaker or the earphone.If the push switch is kept on pushed,the voice can

be heard once each 3 seconds successively.

On the light probe use, a person hasonly to put on the power switch and

tothrow the mode switch to the light mode,then he can hear the note

correspondingto the illuminance at the tip of theprobe. The brighter the

light, thehigher the tone.

Two alkaline dry batteries(size AA)areused to operate this probe. This can

beused continuously about 8 hours as a colour probe and about 12 hours as a

lightprobe without renewing the batteries. If the batteries' voltage down

to belowthe usable level,the voice "renew the batteries"can be heard from

the probe.

6. The test result

In order to ascertain the accuracy of this colour probe,about 300 samples

weremeasured. These are papers, clothes, plastics, leathers, furnitures,

foods,drinks, flowers, leaves and so on. The result shows that although the

accuracyreached to about 95% on the objects with higher chromaticity and

higher lightness, it reduced to a lower figure on the objects with lower

chromaticity andlower lightness.Liquids showed lower accuracy than solids

did.The directivity as a light probe was measured.It has a relatively sharp

directivity of �13 degrees. This means that input light intensity into this

probedecreases to a half value if its pointing direction is deflected 13

degree fromthe initially pointed light source.

7. Evaluation by trial use

This probe is now being used by the visually handicapped persons in order

to beevaluated. According to some users, it will be usefull for

coordinating theclothes being put on in their daily life, for example, a

blouse and skirt, ashirt and trousers, socks and shoes, a shirt and necktie

etc.. Other personsmentioned that after washing many pairs of socks at a

time,it served to findout which two are a pair.It will also be used

effectively at the school for theblind when a teacher gives a lesson in

arts and the blind pupils observe notonly the objects in lesson, but also

everything around them such as flowers,leaves and trees in the garden. In

the office, for example, it can be used tofind a specific file among the

many different coloured files. The light probecan tell the blind persons

the direction of the lamp and if it is on or off.

8. Conclusions

A pocketable colour probe has been realized which discriminates

thirteendifferent colours and tells the operator the name of the colour.

It, moreover,tells him or her the intensity of the light by the tone of

sound.From the viewpoint of accuracy, much effort will have to be made

hereafter.Itwill also be necessary to discriminate dark,vague and

intermidiate colours.

9. Acknowledgment

The auther wishes to thank the many visually handicappedpersons and their

instructors who gave him valuable proposals and informations.He also thanks

to J. Miyajima, k. Muroga and T. Kudo for designing and constructing this

probe.

10. References

(l) F. Furuno : Colour Discriminating Apparatus for the Blind,

IstInternationalConference on Computers for Handicapped Persons, Wien,134-

143,1989

(2) F. Furuno : Colour discriminating apparatus for the blind, Journalof

Microcomputer Applications, 13, 155-164, 1990